## Slide 1
RADIATION-RESISTANT MATERIALS AND SENSORS FOR MAGNETIC DIAGNOSTICS
S. Kulikov 1, I. Bolshakova 2, M. Bulavin 1, I. Vasil’evskii 4, A. Vasyliev 2, V. Chekanov 4, N. Kargin 3, R. Konopleva 4, T. Kuech 5, M. Strikhanov 3, F. Shurygin 2
International Workshop on
Biophysics and Materials at NICA
December 12-13, 2016
Joint Institute for Nuclear Research
Lviv Polytechnic National University
National Research Nuclear University “MEPhI”
Petersburg Nuclear Physics Institute
University of Wisconsin–Madison
1
2
3
4
5

## Slide 2
The materials resistant to the fast neutrons for sensor electronics are intended for the magnetic diagnostics
of charged-particle accelerators and nuclear fusion reactors
Frank Laboratory of Neutron Physics Joint Institute for Nuclear Research Dubna, Russia S. Kulikov, M. Bulavin | Organization and support of experiments on the neutron irradiation of sensors in the channel #3 of IBR-2 reactor
Magnetic Sensors laboratory Lviv Polytechnic National University Lviv, Ukraine I. Bolshakova, A. Vasyliev, F. Shurygin | Development, production, installation in IBR-2 the instrumentations for sensors materials irradiation
Institute of Functional Nuclear Electronics National Research Nuclear University “MEPhI” Moscow, Russia I. Vasil’evskii, M. Strikhanov, N. Kargin | Molecular beam epitaxy method for the semiconductor thick films Vacuum thermal deposition method for the nano-thickness metal films.
Department of Chemical and Biological Engineering University of Wisconsin–Madison Madison, Wisconsin, USA T. Kuech | Metal-organic chemical vapor deposition method for the semiconductor thik films Electron beam evaporation deposition method for the nano-thickness metal films.
Condensed Matter Division Petersburg Nuclear Physics Institute Gatchina, Russia V. Chekanov, R. Konopleva | Experiments on the neutron irradiation of sensors in the channel #V13 of WWR-M reactor
Our international collaboration

## Slide 3: IBR-2 Reactor
Average power, MW | 2
Fuel | PuO2
Number of fuel assemblies | 69
Maximum burnup, % | 9
Pulse repetition rate, Hz | 5; 10
Pulse half-width, µs: fast neutrons thermal neutrons | 240 340
Rotation rate, rev/min: main reflector auxiliary reflector | 600 300
MMR and АМR material | nickel + steel
MR service life, hours | 55000
Background, % | 7.5
Thermal neutron flux density from the surface of the mode- rator: - time average - burst maximum | ~ 1013 n/cm²·s ~ 1016 n/cm²·s
Reactor operation for physics experiments, hr/year ~2500

## Slide 4
Instruments at the IBR-2 reactor
Notes: 14 spectrometers of different types, a neutron imaging facility, two facilities for neutron nuclear physics investigation and a neutron activation facility
4

## Slide 5
IBR-2 horizontal channels diagram
Irradiation facility for channel #3
Diagram of irradiation facility placement in channel #3
1 | – | base part
2 | – | transport beam
3 | – | samples holder
4 | – | samples
5 | – | rail way
1 | – | samples holder
2 | – | transport beam
3 | – | samples
4, 5 | – | biological shields
6 | – | irradiation facility base
7 | – | water moderator
8 | – | shutter of biological shield
9, 10 | – | the stops
11 | – | rail way
A | – | reactor active zone
X | – | distance from moderator surface till samples
Beam #3 of the IBR-2 reactor and irradiation facility

## Slide 6: Parameters of work and neutron spectrum
Time of the reactor cycle – 11 days
Neutron range: from 10-7 to 10 MeV
Maximum Sample size 200 x 400 mm2
Maximal fluence (1 cycle) for neutron with energies more than 1 МeV – 1018 n/сm2
Maximal absorbed dose per cycle at a distance of a sample from the moderator of 40 – 100 МGy (108 J/кg); at a distance 0.3 м – 40 МGy.
Minimal neutron fluence at energy more than 1 МeV – 1015 n/сm2

## Slide 7: Experiments
mini Fcal, Detector Atlas, LHC
Scintillators, LHC
With LHC, CERN
DD modules
Electronic boards

## Slide 8: Electronic boards for ATLAS, CERN
Electronic boards: G10, FR4, Arlon 85N, poliamid (copton), Rogers 4450В. Fast neutron fluenc (Еn>0.4 MeV) ~ 4·1017 н/см2.

## Slide 9
The heads with Hall sensors samples placed on the irradiation facility for channel #3 of IBR-2 reactor

## Slide 10
Personnel area | Technical area | Reactor area
PC, interface unit, voltmeter | System unit | Sensors and magnetic field source
Structure chart of the measurement system
Method for sensors’ on-line testing in nuclear reactors

## Slide 11
Radiation defects
Nuclear doping
Mutual compensation
Acceptors
Donors
Fast
Thermal
Neutrons
(97,5% of possible reactions)
Semiconductor
Minimal sensor parameter drift
(subject to correction with
dedicated electronic methods)
Methods for stabilizing the semiconductor sensor parameters
Chemical doping of semiconductor materials (InSb, InAs) with a complex of doping impurities (donor, isovalent, rare-earth ones) up to the optimal initial concentration of free charge carriers.
Radiation modification – preliminary introduction of a certain number of radiation defects.
Radiation & physical processes in In-containing semiconductor materials under neutron irradiation

## Slide 12
What occurs in sensor’s semiconductor material in the course of irradiation?
In sensors based on InSb semiconductor material, radiation resistance is ensured by two mutually compensating processes:
introduction of acceptor-type radiation defects with fast neutrons
introduction of donors (Sn) by the reactions of transmutation doping on thermal neut-rons according to reaction: (basic reaction: 97,5% of all possible reactions)
For optimal composition of InSb semiconductor material, due to the balance between these two mechanisms and taking account of the ratio between fast and thermal neutrons in flux, sensor parameter drift during irradiation can be minimized, and then subjected to the correc-tion with dedicated electronic methods.
In InAs-based sensors, radiation defects of mostly donor type are generated, and radiation resistance is ensured due to Fermi level pinning in conduction band when achieving optimal donor concentration by two mutually complementing processes:
introduction of donor-type radiation defects when Frenkel pairs (atom in interstice and vacancy) are generated with fast neutrons
introduction of donors (Sn) due to transmutation doping according to reaction:
Radiation & physical processes in sensors’ semiconductor materials during neutron irradiation

## Slide 13
Ex-vessel sensors are planned
to be located in 20 areas
Estimate of neutron flux in the areas of magnetic field sensors ex-vessel location*
* Based on: ITER_D_2F6S7Y v1.2 - Nuclear Environment at Magnetic Sensor Positions (https://user.iter.org/?uid=2F6S7Y)
Neutron fluence in areas 1-9 and 14-20 over ITER lifetime F = (1016÷1017) n∙cm-2
Neutron fluence in areas 10-13 over ITER lifetime F = 1018 n∙cm-2
ITER: ex-vessel Hall sensor locations and neutron fluxes

## Slide 14
Testing results of sensors based on InSb/i-GaAs heterostructure when exposed to reactor neutrons:
curve 1 – radiation-resistant sensors;
curve 2 – industrial sensors.
Comparison with sensors
produced by well-known companies:
LakeShore Cryotronics Inc. (USA)
MetroLab Technology SA (Switzerland)
Testing the radiation resistance of InSb sensors
InSb-based sensors are operable up to neutron fluences of F = 5·1018 cm-2 | 
F = 1015 cm-2 → ΔS/S = 0.04% F = 1016 cm-2 → ΔS/S = 0.08% | F = 1017 cm-2 → ΔS/S = 5% F = 1018 cm-2 → ΔS/S = 10%

## Slide 15
Initial concentration of charge carriers:
curve 1 – n = 1.7∙1018 cm-3
curve 2 – n = 2.4∙1018 cm -3
Dependence of sensitivity of the InAs/i-GaAs
Hall sensors from the neutron fluence

## Slide 16
Dependence of sensitivity of the gold-based
Hall sensors from the neutron fluence
Fabrication method:
curve 1 – electron-beam evaporation
curve 2 – vacuum thermal deposition

## Slide 17
Concept of instrumentation construction is based on:
application of radiation resistant semiconductor magnetic field sensors;
methods of in-situ self-diagnostics, periodic calibration and correction of signals without sensors’ removal.
Instrumentation includes:
3D probe with integrated magnetic measuring transducers containing radiation resistant Hall sensors and test microsolenoids;
control electronics unit;
personal computer;
interfaces and signal transmission lines;
software.
Parameters:
Magnetic induction measurement range ………. (1∙10-4÷5) T
Accuracy of measurement ………………………. 0.1 %
Resolution …………………………………………. 0.1 mT
Magnetic measuring instrumentation for radiation operating environment

## Slide 18
HG
HG
Coil
Probes with 3D integrated magnetic transducers

## Slide 19
The existing system of ITER magnetic diagnostics is based on conventional pick-up coils with integrators, which do not ensure the required accuracy of measurements:
insufficient accuracy of steady-state magnetic field measurement at long pulses,
inertness of integrators at high-frequency measurements,
radiation-induced parasitic effects (RIC, RIEMF, RIED).
to supplement the existing ITER magnetic diagnostics system with galvanomagnetic instrumentation based on radiation-resistant semiconductor Hall sensors, capable of measuring the steady-state magnetic field at long pulses with high precision (error less than 0,1%) under irradiation with high neutron fluences
The proposal :
Problems in tokamak
magnetic diagnostics
Application of radiation resistant sensors and
instrumentation for magnetic diagnostics in fusion reactors

## Slide 20
3D-probe location in Tore Supra
Results of magnetic field measurement
at Tore Supra in experiment # 34085
Magnetic field measurement in Tore Supra (France)

## Slide 21
JET cross-section
Magnetic field measurement results at JET in experiment # 64416
Magnetic field measurement in JET (UK)

## Slide 22
Main achievements:
Radiation-resistant semiconductor magnetic field Hall sensors that can operate in ITER-relevant conditions (neutron fluences F > 1018 n·cm-2) have been created
Magnetic-measuring instrumentation with 3D Hall sensors have been developed and successfully tested in TORE SUPRA and JET tokamaks
Long-term and stable operation of 3D Hall sensors is confirmed by their successful use in JET tokamak over 5 year period (2009-2014)
New methods for tokamaks’ magnetic field measurement have been suggested and patented: FR 2 887 991 B1, GB2427700, Application №10858732.0. The methods allow us to meet ITER requirements concerning magnetic diagnostics
Sensors’ long-term stability (± 0.1 %) in probe # 52 during 8000 JET pulses (2009-2014).
Histogram of calibration data from all 18 Hall sensors over their 5-year operation on JET
Experimental evaluation of stable long-term operation of semiconductor sensors in ITER-relevant environment
July 2014

## Slide 23
Magnetic field mapper of cyclotron magnet for
National Institute of Radiological Sciences (NIRS), Japan

## Slide 24
Magnetic systems of the NICA/MPD
accelerator complex

## Slide 25
Accuracy of measuring electronic channels .............................................. 0.01%
Accuracy of signal readout from each sensor ......................................... (0.04-0.08)%
Magnetic field resolution ……………. 0.1 mT
Measurement frequency ……………... 1 per 100 ms
Range of measured magnetic field … (0.05-1.5) T
3D magnetometer MSL-3DM1 for NICA collider

## Slide 26
Hall sensors in MSL-3DM1 magnetometer

## Slide 27
Model Manufacturer | MSL-3DM1 Magnetic Sensor Laboratory Lviv, Ukraine | THM 1176 MetroLab, Switzerland | Model 460 LakeShore, USA
Measuring accuracy | 0,1% | 1% | 0,35%
Resolution | 0,1 mT | 3 mT | 0,1 mT
Sensing element dimensions for one sensor | 0,2 х 0,2 mm2 | 0,15 х 0,15 mm2 | 0,5 х 1,0 mm2
3D-probe active volume dimensions | 1,8 х 1,8 х 1,8 mm3 | 6,0 х 3,4 х 6,0 mm3 | 4,0 х 4,0 х 3,6 mm3
3D-probe dimensions | 40 х 19,7 х 13,6 mm3 | 113 х 16 х 10 mm3 | 127 х 6 х 6 mm3
Operating temperature range | 5 ÷ 25 ºС | 0 ÷ 40 ºС | 10 ÷ 40 ºС
Technical features of MSL-3DM1 magnetometer
Comparison with contemporary other Hall 3D magnetometers
Measurement accuracy 0.1% with MSL-3DM1
is better than 0.35% for Model 460 (LakeShore) and 1% for THM1176 (MetroLab)

## Slide 28
Requirement specification
Magnetic field induction ………………… 0.5 T
Magnetic field homogeneity ……………. 4·10-4 T
Accuracy of relative measurements …… ≤ 0.1%
Radiation load …………………………… to be defined
:
Multi-Purpose Detector (MPD)
Mapping of MPD detector magnetic field
Proposal for future collaboration

## Slide 29
Thank you for your attention!